## Slide 1: Lecture 16: Nuclear Imaging: System & Applications
Dr. Guohua Cao （曹国华） caogh@shanghaitech.edu.cn
School of Biomedical Engineering
ShanghaiTech University
4/2/2023

## Slide 2
Signal-to-Noise Ratio
Dominate source of image noise is due to relatively limited number of total γ-ray events
Counts in each image histogram bin are Poisson distributed
Total number of counts (N) is ~1.0e6
SNR Due to Poisson Counting Statistics Alone
SNR  E[N ] N
 N
N
 N
Total counts depends on activity of radiotracer, sensitivity of gamma camera, acquisition time, and tissue attenuation


## Slide 3: SPECT Overview
One or more gamma cameras are rotated around the subject during data acquisition
Estimates of the 3D tracer distribution are computed with tomographic reconstruction algorithms

## Slide 4: SPECT Acquisition
Just as for X-ray CT, detectors are rotated around the subject, to acquire fully tomographic data sets
Contoured Orbit
Circular Orbit Sinogram
Detectors must be rotated >= 180 deg
Due to close proximity to subject and large mass of detectors rotation time is ~ couple of minutes

## Slide 5: Gamma Camera Configurations
Single Headed Triple Headed
Double Headed
Different clinical applications utilize specific detector configurations
Increasing the number of gamma cameras and closeness to subject leads to increase in image SNR!

## Slide 6: SPECT Scanners
Dedicated Cardiac System
Small Animal Scanner
Multi-purpose System

## Slide 7: Image Reconstruction
System Matrix
Subject Attenuation
Data
Activity Distribution
Subject Compton Scatter
y  PoissonPAx  s
Images are quantitative if subject attenuation and scatter are corrected for!

## Slide 8: Iterative Image Reconstruction

## Slide 9: Linear versus Iterative Reconstruction
Iterative reconstruction algorithms give superior image quality over linear methods for nuclear imaging
Filtered Backprojection (FBP) Iterative
Comparison of Iterative and FBP Reconstructions
Iterative has significantly higher SNR, especially for limited count acquisitions
Iterative is much less prone to artifacts, due to incorporation of system response
Iterative is much more computationally complex due to multiple forward and back-projection operations
FBP is a linear algorithm, while iterative methods are non-linear

## Slide 10: Attenuation Correction
Uncorrected attenuation leads to an underestimation of tracer uptake in reconstructed images
Challenging since attenuation is depth dependent

## Slide 11: Attenuation Correction Methods Chang’s Multiplicative Method

## Slide 12: Measured Attenuation Correction

## Slide 13
Scatter Correction
Scatter causes a low spatial frequency offset to SPECT images if not corrected for

## Slide 14: SPECT-CT Scanners
Allows localization of tracer distribution wrt anatomy
CT images can be used to estimate attenuation at γ- ray energy

## Slide 15: SPECT Radiotracers
Most frequently 99mTc labeled radiotracers

## Slide 16: Myocardial Perfusion Imaging (MPI)
99mTc-sestamibi
Tracer uptake controlled by local blood flow
Exercise increases blood flow (3-5X higher), which increases dilation of normal arteries, but not diseased ones.

## Slide 17: Myocardial Perfusion Imaging (MPI)

## Slide 18: Myocardial Perfusion Imaging

## Slide 19
Bone and Oncology Imaging
99mTc-methyl diphosphonate (MDP)

## Slide 20: PET Overview
Positron (β+) emitting radiotracers are injected into the subject (Always positron emitter!)
The two back-to-back annihilation photons are detected simultaneously by separate detectors
Annihilation photon path is determined by position of the two involved detectors

## Slide 21: PET Overview

## Slide 22: United Imaging Total Body PET

## Slide 23: 联影 uExplorer

## Slide 24: The World‘s First Total-body PET-CT uEXPLORER
https://www.youtube.com/watch?v=qGElYRqP8SA

## Slide 25: Annihilation Coincidence Detection (ACD)
Coincidence Timing Window (6-12 ns)
Annihilation Photons
Detected Coincidence

## Slide 26: Annihilation Coincidence Detection
Detector
Line of Response
Annihilation Photons
ACD significantly restricts the volume that detectors are sensitive to annihilation photons
Only annihilation photons that originate at locations within the tube or line of response (LOR) can be detected
ACD is form of electronic collimation (vs. absorptive for SPECT)

## Slide 27: Coincidence Binning
LORs can be formed between many detector combinations
The location (index) of detectors involved in a coincidence detection determine the sinogram bin
Sinogram
0
1
2
3

## Slide 28: PET Scanners Whole-Body Scanner
Small Animal Scanner
Dedicated Brain Scanner

## Slide 29: PET Detectors
Detector blocks are arranged into detector rings to maximize gamma-ray detection

## Slide 30: Block Detector
Blocks are partially segmented into long crystals (~2 cm)
Spread of scintillation light on PMTs dependent on interaction position
Four PMTs read out scintillation light and position events with Anger Logic

## Slide 31: Scintillators
511 keV gamma-rays require scintillators to have high LACs to be absorbed
Timing resolution (ability separate events in time) is reflected in the decay time (lower decay time=improved resolution=smaller coincidence window)

## Slide 32: Positron Decay + Migration p  n  e    energy
511 keV photon
511 keV photon
Annihilation Reaction
Effective Positron Range
Nucleus
Positron
Positron Path
Positrons will migrate and “bounce” off electrons before undergoing an annihilation reaction
Causes annihilation point to be different from radiotracer location!

## Slide 33: Effect of Positron Migration on Resolution
Positron range depends on emission energy of the positron and material of interaction (FWTM=1-4 mm)
Increasing β+ energy
Increase in β+ range Degradation in resolution

## Slide 34: Photon Non-colinearity
180

R  0.0022  D
Non-colinearity spatial resolution
D= PET detector diameter
θ
D
Non-zero positron momentum during annihilation causes 511 keV photons not to be exactly back-to-back (θ≠180°)
Becomes worse as scanner diameter increases (~2 mm)

## Slide 35: Detector Spatial Resolution
Depends primarily on the face dimensions of the pixelated scintillator elements
Increase face dimension
Degrades spatial resolution
Typically the dominant source of spatial resolution loss for PET scanners (~4 mm)

## Slide 36: 2D and 3D PET
2D PET restricts coincidences axially to direct and cross- planes
3D PET utilizes nearly all available axial detector combinations
3D PET has significantly higher photon efficiency than 2D PET (>=50% increase in reconstructed image SNR)
2D
3D
Cross-Section View

## Slide 37: Image Reconstruction
y  PoissonAPx  s  r
System Matrix
Subject Attenuation
Data
Activity
Distribution
Subject Compton Scatter
Iterative reconstruction is primarily, as in SPECT, but reconstruction problem is different in PET
Randoms

## Slide 38: Coincidence Types
Random
True
Scatter
Scatter

## Slide 39
Randoms (Accidental Coincidences)
Definition: Coincidences that are formed between annihilation photons originating from separate emissions
C
A
 2 S S
ij i j
i
j
between detectors i and j
= coincidence window (sec)
CA = random coincidence rate
ij

Sx = singles rate on detector x (cps)
Randoms cause very low frequency positive bias to reconstructed uptake estimates if not corrected for
Rates of singles (detected γ-rays used to perform ACD) are increased with higher injected activity

## Slide 40
Attenuation Correction
Line of Response
i
j
L
Li
Lj
g(x
)
N
N
µ(x)
Number of Counts
x=0
g(x) = annihilation events at x Px
µ(x) = linear attenuation at x
= probability of photon detection

## Slide 41: PET-CT
Subjects are scanned sequentially
Almost all PET systems sold today are combined with CT
Significantly improves tracer localization, fusion registration accuracy, and improves photon attenuation correction

## Slide 42: PET-MRI
cular
11B
11C
MRI
11A
PET-MRI Scanner
PET
Fusion
Modern systems image subjects simultaneously
Combines very high soft tissue diagnostic sensitivity of MRI with high specificity of PET
May have advantages in brain, cardiac, and focused oncology imaging compared with PET-CT